# MockData1 Validation Test Results

Date: April 2, 2026

Validator: DVaaS (jsonschema-rs + Allotrope ASM schemas, 113 schemas loaded)

# Executive Summary

We ran both files through our validation pipeline. MockData1_2.json correctly failed with 3 schema errors. MockData1_1.json passed schema validation, which revealed an important finding: the official Allotrope JSON schemas are more permissive than expected. They define field types and structure but enforce very few required fields. Issues like missing analyst, value rounding, and empty strings are not caught by schema validation alone. This confirms that multiple validation layers are needed, aligning with your 4-level validation approach.

# MockData1_2.json (Row 2) - INVALID

3 schema errors detected. Our Level 1 validation correctly flagged these:

| Error | Measurement | Root Cause |
| --- | --- | --- |
| 1 | Blood gas (measurement[0]) | pCO2 unit is "nmHg" instead of "mmHg" |
| 2 | Metabolite (measurement[3]) | glucose unit "mg/L" not valid; schema expects "g/L" |
| 3 | Cell counter (measurement[4]) | cell density dilution factor missing "unit" field |

# MockData1_1.json (Row 1) - VALID (Schema Passed)

0 schema errors. The file passes official Allotrope JSON schema validation.

## Why It Passed

The official Allotrope JSON schema (REC/2025/12) is permissive by design. The "analyst" field is optional (not in a required array). Value precision is not enforced. Empty strings are valid. Extra characters in strings pass type validation. The schema validates structure and types, not data quality or business rules.

## Level 2 Issues Found by Manual Comparison

| # | Source Field | Source Value | ASM Value | Issue |
| --- | --- | --- | --- | --- |
| 1 | PO2 | 164.7 | 165 | Value rounded |
| 2 | O2 Saturation | 100 | 90 | Value changed |
| 3 | Chem Cartridge Lot | 25099058 | 25099058-- | Extra characters appended |
| 4 | Chem Card Lot | 25094029 | (empty) | Value dropped entirely |

# Gap Analysis: Schema vs Customer Expectations

Your Level 1 validation includes both JSON schema validation AND ASM tag validation via your custom Python package. Our current DVaaS performs the JSON schema part. The additional rules your package enforces are not yet in our service.

## What the Schema Enforces

Document structure and hierarchy

Required fields: measurement identifier, measurement time

Field types (string, number, object, array)

Detector schema conformance

$asm.manifest presence

## What the Schema Does NOT Enforce

analyst as a required field (optional in schema)

Value precision or accuracy

Non-empty string constraints

Specific unit values beyond detector schemas

Business rules

## Our Plan to Close the Gap

| Rule | Source | Priority | Status |
| --- | --- | --- | --- |
| $asm.type attribute enforcement | Validation workflow / sdc-asm-to-rdf | High | Planned |
| $asm.pattern enforcement (value datum = value + unit) | Validation workflow / sdc-asm-to-rdf | High | Planned |
| Required analyst field | Building Principles | High | Planned |
| Non-empty string validation | MockData1 test findings | High | Planned |
| Calculated data traceability nesting | Traceability example | High | In progress |
| Every source term mapped (no silent drops) | Mapping rules | High | Implemented (field_mapping) |
| Unit validation against ontology | Building Principles | Medium | Planned |
| HarmonizingBaseTerms integration | Reference Excel | Medium | Planned |
| Configurable L3 vs L4 rules | Mapping email | Medium | Planned |
| Validation level labels on findings | Validation workflow | Medium | Planned |

# Validation Coverage Summary

| Issue Type | Schema (Level 1) | Data Integrity (Level 2) |
| --- | --- | --- |
| Wrong unit (nmHg) | Caught | Would catch |
| Invalid unit for field (mg/L) | Caught | Would catch |
| Missing unit field | Caught | Would catch |
| Value rounding (164.7 to 165) | Not caught | Would catch |
| Value changed (100 to 90) | Not caught | Would catch |
| Extra characters appended | Not caught | Would catch |
| Value dropped to empty | Not caught | Would catch |
| Missing analyst field | Not caught (optional) | N/A - add as rule |

# Timestamp and Date Format Analysis

| Aspect | Row 1 | Row 2 | Notes |
| --- | --- | --- | --- |
| Source timestamp | 2/7/2025 10:35 | 2/7/2025 10:35:00.123 | Row 1 missing seconds; Row 2 has milliseconds |
| ASM timestamp | 2025-02-07T17:35:00+00:00 | 2025-02-07T17:35:00.123+00:00 | Both correctly offset +7 hours (MST to UTC) |
| Seconds handling | Assumed :00 | Preserved .123 | Converter fills missing seconds with zero |
| Date interpretation | February 7 (US) | February 7 (US) | Ambiguous without timezone context |

Timezone conversion is correct - 10:35 MST + 7 hours = 17:35 UTC.

Date format ambiguity - 2/7/2025 is M/D/YYYY (US) or D/M/YYYY (European). Without timezone or locale context, this is ambiguous. This is the [META] issue from your review. Our instrument config supports "timezone": "America/Denver" to resolve this.

Millisecond precision in Row 2 - Source has 10:35:00.123 which is unusual for this instrument. ASM correctly preserves milliseconds.

Missing seconds in Row 1 - Source has 10:35 without seconds. Converter assumes :00. Level 2 comparison should note this inference.

# Recommendations

1. We are reviewing the materials you shared - Thank you for providing the validation workflow, mapping rules, traceability example, HarmonizingBaseTerms Excel, and Building Principles. We are also reviewing your $asm tag validation approach (sdc-asm-to-rdf). We will use these to add supplementary validation rules to DVaaS beyond the base JSON schema.

2. Calculated data traceability - Your traceability example confirms the pattern: data source aggregate document nested inside each calculated data document. We are updating our converter to match.

3. Two-layer validation rules - We agree with separating Allotrope-approved rules (Level 3) from organization-specific rules (Level 4). We will implement both as configurable rule sets in DVaaS.

4. Every source term must be mapped - Our converters follow this principle (all 40 CSV columns mapped for Nova FLEX2). Our Custom Converter Requirements document mandates this for all customer-built converters.

5. Test Level 2 via conversion pipeline - Upload MockData1.csv through our "Convert Instrument File" tab to see Data Integrity Verification in action.

6. HarmonizingBaseTerms integration - We will review the Excel file and incorporate standardized terms as converter configuration input.

## Timezone Verification

Source: 2/7/2025 10:35 (MST). ASM: 2025-02-07T17:35:00+00:00 (UTC). 7-hour offset correct. Our instrument config supports location.timezone for this purpose.