# AWS Response to Merck ASM Conversion Review

In response to: 2026-03-17 AWS ASM Conversion Review & Merck ASM Basic Design Principles v33

# Responses to Questions

## Question 1: How can the platform capture instrument-specific metadata?

The instrument config file (JSON) is the right vehicle. We will extend the schema to include metadata fields that capture the information currently derived from the mount path: geography, site, UNC path, and timezone.

The timezone field solves the issue where the July file (European) was interpreted with US date format. The UNC path field provides traceability back to the source file server. The instrument config is created once per instrument and reused for every file.

For the future Greengrass edge client deployment, the mount path metadata would be captured automatically since the edge client knows which instrument folder it is watching.

## Question 2: Can the AI agent provide a confidence level for each mapping?

Yes, but with an important distinction between conversion routes:

Custom converters (Route A): Confidence is 100% by design. Every field mapping is deterministic. The Data Integrity Verification report proves this field-by-field.

Allotropy library (Route B): Confidence is 100% for supported instruments. Rule-based parsers, no AI inference.

ATaaS / AI-powered (Route C): This is where per-mapping confidence levels are valuable. We can enhance ATaaS to provide High/Medium/Low confidence per field mapping.

We recommend custom converters or allotropy for production use where confidence is deterministic. ATaaS is best suited for initial analysis of unknown formats with human-in-the-loop review.

## Question 3: Which validation levels are we using today?

| Merck Level | Description | Our Equivalent | Status |
| --- | --- | --- | --- |
| Level 1 | ASM schema validation (JSON + ASM tags) | DVaaS with allotropy validation script | Implemented |
| Level 2 | Data quality assessment (ASM vs raw data) | Data Integrity Verification (field_mapping) | Work in progress — Nova FLEX2 done, rolling out |
| Level 3 | Allotrope ASM building principles | Not yet implemented | Planned — add as DVaaS rules |
| Level 4 | Merck-specific principles | Not yet implemented | Planned — configurable per-customer rules |

Our Data Integrity Verification directly addresses your Level 2 pain point. You noted that Level 2 requires "significant instrument-specific coding." Our approach solves this: the converter emits the field_mapping during conversion, so verification is automatic. No TOSCA database, no SQL queries, no instrument-specific scripts.

We will update DVaaS to tag each finding with the corresponding validation level (L1/L2/L3/L4) so the human-in-the-loop reviewer can prioritize accordingly.

# Remediation Plan

## Issue Tracker

| # | Category | Issue | Severity | Status |
| --- | --- | --- | --- | --- |
| 1 | [META] | Measurement time missing timezone/geography context | Medium | Open |
| 2 | [ERROR] | Sample type uses invalid ASM term | High | Open |
| 3 | [ERROR] | Skipped processed data: total cell count, viable cell count, cell density dilution factor | High | Open |
| 4 | [ERROR] | Cell density dilution factor not nested under data processing document | Medium | Open |
| 5 | [ERROR][RULE] | Calculated fields were ignored | High | Fixed |
| 6 | [ERROR][RULE] | Custom fields were ignored | High | Fixed |
| 7 | [CONFIG] | Missing Data System Document | High | Partially fixed — needs UNC path |
| 8 | [CONFIG] | Missing Device Document with device type | High | Fixed |
| 9 | [META] | No traceability to original source file (UNC path) | Medium | Open |
| 10 | [RULE] | Custom fields have manually added units not in source | Medium | Open |
| 11 | [RULE] | Missing @index for array ordering | Low | Open |
| 12 | [RULE] | Calculated data missing per-calculation data source linkage | Medium | Open |
| 13 | [RULE] | No processed data aggregate document layer | Medium | Open |

## Already Fixed Since Review

Issue 5 — Calculated fields: Converter now includes calculated data aggregate document with temperature-corrected pH, PO2, PCO2, and bicarbonate.

Issue 6 — Custom fields: Converter now maps all custom fields to custom information aggregate document including lot numbers, flow times, vessel pressure, dilution ratio, and all string metadata.

Issue 8 — Device Document: Converter includes device system document with device identifier, product manufacturer, and device type = "solution analyzer".

## Phase 1 — High Priority

| # | Issue | Effort | Change |
| --- | --- | --- | --- |
| 1 | Timezone in instrument config | 1 hour | Add location.timezone to config, use in timestamp parsing |
| 2 | Invalid sample type term | 30 min | Map to valid ASM term per schema |
| 9 | UNC path traceability | 1 hour | Read from instrument config, write to data system document |
| 10 | Remove manually added units from custom fields | 1 hour | Only include units when present in source data |
| 12 | Per-calculation data source linkage | 1 hour | Nest data source aggregate inside each calculated data document |

## Phase 2 — Medium Priority

| # | Issue | Effort | Change |
| --- | --- | --- | --- |
| 3 | Processed data (cell counts, dilution factor) | 2 hours | Add processed data aggregate document |
| 4 | Dilution factor nesting | 30 min | Nest under data processing document |
| 11 | @index for array ordering | 1 hour | Add @index based on source order or timestamp |
| 13 | Processed data layer | 2 hours | Distinguish raw measurement vs instrument transformation |

## Phase 3 — Validation Enhancement

| Item | Effort | Change |
| --- | --- | --- |
| DVaaS Level 3 rules (Allotrope principles) | 4 hours | Add design principle checks to validation |
| DVaaS Level 4 rules (Merck-specific) | 4 hours | Add configurable customer-specific rules |
| Validation level labels on findings | 2 hours | Tag each error/warning with L1/L2/L3/L4 |
| ATaaS confidence scores per mapping | 4 hours | Add per-field confidence to AI-generated mappings |

# Data Integrity Verification — Addressing Level 2

Your Level 2 validation (TOSCA-based source-to-ASM comparison) requires significant per-instrument coding. Our Data Integrity Verification provides the same assurance with a fundamentally different approach:

| Aspect | Merck Level 2 (TOSCA) | AWS Data Integrity Verification |
| --- | --- | --- |
| Approach | Post-conversion comparison via SQL | Converter emits field_mapping during conversion |
| Per-instrument effort | Significant — SQLite, SQL queries, tolerances | Zero — field_mapping built into converter |
| Tolerance handling | SQL-based rounding and TOSCA parameters | Exact match — no rounding needed |
| Coverage | Depends on SQL queries written | Every field the converter touches |
| Output | TOSCA report | Dashboard table + API response |
| Maintenance | Update SQL when converter changes | Automatic — updates with converter |

Key advantage: Because the converter emits the mapping at conversion time, there is no separate comparison step, no database, and no instrument-specific SQL. The integrity proof is a byproduct of the conversion itself.

Current status: Implemented for Nova FLEX2 (all 40 CSV columns). Required for all custom converters per our requirements spec. Planned for allotropy-based conversions.

# Alignment with Merck Design Principles

## What We Align On Today

| Principle | Our Implementation | Status |
| --- | --- | --- |
| Instrument hardcoded terms | Instrument config JSON file | Aligned — needs UNC path and timezone |
| Data hierarchical structure | Follow solution-analyzer schema | Aligned |
| ASM terms lowercase with spaces | Terms match schema (pO2, pCO2, pH, osmolality) | Aligned |
| Custom fields 1:1 transfer | Custom information aggregate document | Aligned — need to remove manual units |
| Calculated fields with traceability | Calculated data aggregate document | Aligned — need to fix nesting |
| Manufacturer in lowercase | "nova biomedical" | Aligned |
| Source file traceability | Data system document with file name, converter info | Partial — need UNC path |

## What We Need to Add

| Principle | Gap | Plan |
| --- | --- | --- |
| Processed data aggregate document | Not implemented | Phase 2 |
| @index for array ordering | Not implemented | Phase 2 |
| HarmonizingBaseTerms integration | Not formalized | Accept as input alongside instrument config |
| Custom field units only from source | We add units manually | Phase 1 |
| Statistic datum role | Not implemented | Future |

# Next Steps

1. Share this response with Merck for review

2. Phase 1: Implement high-priority fixes (timezone, sample type, UNC path, custom field units, data source nesting) — estimated 5 hours

3. Phase 2: Implement medium-priority fixes (processed data, @index, dilution factor nesting) — estimated 6 hours

4. Phase 3: Validation enhancement (L3/L4 rules, level labels, confidence scores) — estimated 14 hours

5. Ongoing: Accept and integrate HarmonizingBaseTerms spreadsheet as converter configuration input

Built with: AWS Lambda, AWS Bedrock Claude, Allotropy Library, React, AWS CDK